## Slide 1: Streamgaging using surface-velocity sensors (RQ-30s) Presented by Scott Grzyb, USGS to the TWDB FLOG/FIM Subcommittee 5 April 2024
U.S. Department of the Interior
U.S. Geological Survey

## Slide 2: Objectives
Densified forecasting requires densification of stream gaging
Assess if newer technology can achieve comparable results with less effort

## Slide 3: Study Area 80 TxDOT gages with RQ-30 radar sensors

## Slide 4
Traditional USGS Discharge Determination
Flowtracker Handheld ADCP
Acoustic Doppler Current Profiler
USGS Stage-Discharge Rating
Measure Stage Height
Use Rating Curve to
get Discharge

## Slide 5
Measurement of Water Elevation and Velocity using RQ-30 Radar Gage
Discharge = Area * Mean Velocity

## Slide 6
Computation of Discharge using RQ-30 Radar Gauge
A
V
Q

## Slide 7: Cross Section Surveying 08041940-Green Pond Gully at FM 365 nr Fannett, TX
Channel geometry in-line with the velocity sensor
Low chord
Upper bridge deck / road elevation
Sensor elevation

## Slide 9
K-Factor Profile
Cross-section

## Slide 10: Gage Height Accuracy

## Slide 11: Gage Height Peak Verification

## Slide 12
Gage Height Assessment
Gage Heights for 53 stations
No resets made for the life of the gage
+/- .05 ft accuracy
+/- 0.05 ft

## Slide 13: Velocity vs Learned Velocity
Date/time
Velocity ( ft/s)

## Slide 14: Point Velocity (surface)
Date/time
Velocity ( ft/s)

## Slide 15: Learned Point Velocity (surface)
Date/time
Velocity ( ft/s)

## Slide 16: Early discharge observations
November 2021
USGS Peak Discharge – 1,539 cfs
RQ-30 Peak Discharge – 1,500 cfs
July 2021
USGS Peak Discharge – 665 cfs
RQ-30 Peak Discharge – 680 cfs
Date
Discharge ( cfs )
08167000 – Guadalupe Rv at Comfort, TX

## Slide 17: Discharge observations
Date/time
Discharge ( cfs )
08111051 - Carters Ck nr College Station, TX
Measurements
RQ-30 Discharge

## Slide 18: Discharge observations
Date/time
Discharge ( cfs )
08111110 - New Year Ck at FM 1155 nr Chappell Hill, TX
Measurements
USGS Discharge
RQ-30 Discharge

## Slide 19: Discharge Measurement Comparisons
Date/time
08189718-Chiltipin ck at us 77 nr Sinton TX
Discharge ( cfs )
Measurements
USGS Rated Discharge
RQ-30 Learned Discharge
Measurements are 29% and 28% lower than USGS rated discharge.

## Slide 20: Q-Commander Calibration

## Slide 21: 081111110 – New Year Ck at FM 1155 nr Chappel Hill, TX
Number of Calibration Points

## Slide 22: Targeted Approach
Measure a minimum of 2 discharges at different channel geometries.
Calibrate the RQ-30
Once calibrated, make minimal measurements to verify accuracy

## Slide 23: Discharge observations – Pre-Calibration
Date/time
Discharge ( cfs )
08111110 - New Year Ck at FM 1155 nr Chappell Hill, TX
Measurements
USGS Discharge
RQ-30 Discharge
Peak 700 cfs

## Slide 24: Discharge observations – Post Calibration
Date/time
Discharge ( cfs )
08111110 - New Year Ck at FM 1155 nr Chappell Hill, TX
Measurement
USGS Discharge
RQ-30 Discharge
Peak 600 cfs

## Slide 25
Measurement Number
081111110 – New Year Ck at FM 1155 nr Chappel Hill, TX

## Slide 26: Targeted Measurement Approach Exercise
Students independently select 2 measurements they think represents channel geometry
Calibration is performed through Q-commander (RQ-30 software)
Results are compared with classmates
(from USGS Velocimetry training class)

## Slide 27: Externalizing K-Factor Calculation
Existing USGS app Area Comp 3 Is being modified to calculate K-factor - the coefficient to convert surface velocity to mean channel velocity

## Slide 28: Externalizing Learned Velocity
Learned velocity is determined within the RQ-30
Currently no way to transfer learned velocity
Future learned velocity profile calculation being developed for integration into Area Comp 3

## Slide 29: Relationship between averaged cells
Distance from Left Bank
Depth

## Slide 30: Relationship between averaged cells
Distance from Left Bank
Depth

## Slide 31: K-Factor Methods Compared
Q-Commander
ADCP Deconstruction

## Slide 32
USGS Stage-Discharge Rating
Q-Commander
ADCP Deconstruction
Numerous Measurements
2 - 4 Measurements (Using Targeted Approach)
1 - 2 Measurements
No need for roughness

## Slide 33: Effectiveness in dynamic channel conditions
July 2021
February 2022
Seasonal changes

## Slide 34: Effectiveness in dynamic channel conditions cont.

## Slide 35: Cross Section Surveys 08111051-Carters Ck at FM 60 nr College Station
Main channel shows approximately 30% reduction in area at low flows following the application of riprap on the right bank.

## Slide 36: Observations After Rectification

## Slide 37: Tidally Affected Gages
08031020 – Cole Ck at I-10 nr Orange, TX
Stage
Velocity

## Slide 38: Backwater Conditions Cont. Relation between Stage Height and Velocity (Tehuacana Creek nr Waco)
Discharge Stg/Q
Discharge by velocity

## Slide 39: Data Observations – Velocity
08164150 - W Navidad Rv at I-10 nr Schulenburg, TX
Date/time
Stage ( ft )

## Slide 40: Data Observations – Velocity
08164150 - W Navidad Rv at I-10 nr Schulenburg, TX
Date/time
Stage ( ft )

## Slide 41: Data Observations – Initial Velocity Wave
08111110 - New Year Ck at FM 1155 nr Chappell Hill, TX
Chart: Peak Initial Pulse Velocity vs Water Level
| 2.11 | 2.11 | 2.2400000000000002 | 2.36 | 2.5499999999999998 | 2.78 | 2.09 | 2.0299999999999998
 | 156.78 | 156.44 | 162.02 | 160.1 | 162.84 | 164.1 | 156.15 | 156.44
Stage (ft)
Peak velocity of initial pulse

## Slide 42: Data Observations – Initial Velocity Wave
08111110 - New Year Ck at FM 1155 nr Chappell Hill, TX
6 hours
19 hours
Subsequent rain events

## Slide 43
 | Questions?